[image: image1.png]



[image: image20.png]CYBERTEK vt

Test & Measurement






Optical Isolated Voltage Probe

OP6031A (3V~2500V/150MHz)

OP6033A (3V~2500V/350MHz)

OP6035A (3V~2500V/500MHz)

[image: image21.png]



Shenzhen Zhiyong Electronics Co., Ltd

Preface

First of all, thank you for purchasing our products, this instruction manual is the description about the function, usage, operation attention points, etc. Before use, please read the instructions carefully and use correctly. 

Manual annotation will use the following symbols to distinguish.

 SHAPE  \* MERGEFORMAT 

[image: image2]




To the safely use the machine, you must abide by the following safety precautions strictly. The violation against the manual is likely to damage the protective function of the machine. In addition, the company is not responsible for any safety problem caused by the violation of matters needing attention in operation.


Please be careful to the danger of electric shock and pay attention to highest input voltage.

Do not operate in wet or combustible conditions.

Make sure the circuit under test is turned off before access it to the probe.

Turn off the circuit after the measurement, and then remove the probe.

When BNC cables are connected to the oscilloscope or other devices, ensure the BNC terminal is well grounded.

Check the probe skin and probe lead regularly. If there is any breakage, stop using it immediately.
OP6000 series products briefing
	Model
	Maximum Input Voltage
	Bandwidth
	Attenuation Ratio

	OP6031A
	2500Vpk(2000X)
	150MHz
	20X(standard)

2X/500X/2000X (optional)

	OP6033A
	2500Vpk(2000X)
	350MHz
	20X (standard)

2X/500X/2000X (optional)

	OP6035A
	2500Vpk(2000X)
	500MHz
	20X (standard)

2X/500X/2000X (optional)


Introduction

OP6000A seriesis the latest optical isolated voltage probe with remarkably high CMRR. The CMRR of traditional differential probe decreases fast in high-frequency range, as a result, measuring the small voltage signal waveform (e.g., the driving voltage when measuring the upper MOSFET of the half-bridge circuit) under high CM interference voltage accurately become a extremely hard task. OP6000A seriesis applay optical isolation technologies and gains remarkably high CMRR in all working bandwidth, helping our customers to deal with these kinds of challenging measurement with low cost.  
Product Characteristics:

Bandwidth over 500MHz

Isolation voltage over 60kV

High CMRR
Can be calibrated and zeroed online without disconnecting from the tested equipment
It will automatically go into sleep mode 5 seconds after the receiving end is powered off.
High precision in wide temperature range
Able to measure gate voltage drive of Si/SiC/GaN and other power devices 
Able to measure high-frequency current of Si/SiC/GaN and other power devices with coaxial shunt CSD series.
2 Chargeable batteries that can be replaced to keep the probe work continuously
Application

OP6000A series can be widely applied in the R&D, debugging or maintenance of switching power supply, motor driver, new energy inverter, converter, LED power supply, household appliances and other electrical power devices. 

Floating signal test

The gate driving signal measurement of upper bridge MOSFET can also be applied to the measurement of high-speed driving waveform of SiC and GaN.

Small signal measurement of differential mode under high common mode voltage

Products and Accessories

█ Probe

[image: image3.jpg]vV Y

A AdAET
02X NAA0ETF
005X AdA00SF

0002x | ¥dA00SZF
OolLVY JONWVY

3903¥d 3OVLIT0A TVIIlLdO

ZHIN0OS
vG€09dO

MALAATAD

S

O &

N J
oMENB:< S

7

«__" )

.

Jlaseq] [J Jamod (O

390¥dd dOV110A 1VIIlLdO

1\

AHLAHIAD

N\

/

® ©®




Input connector:  MMCX interface. Attenuator for connecting products. Standard CK-AT20XB (20X), optional CK-AT2XB (2X), CK-AT500XB (500X), CK-AT2000XB (2000X), can be selected according to voltage measurement needs.
BATTERY indicator: lighted green under normal condition; lighted red under insufficient power supply, please change the battery in time.

Optical fiber: connecting the front-end light emitter and rear-end light receiver, 1.5m in length.
 

Please refer to the tips below when using:
Do not put heavy objects on the optical fiber to avoid any stress on it.

Do not squeeze or curl the optical fiber and do not curl the optical fiber into a loop diameter less than 10cm.

Do not twist or knot on optical fiber and do not pull it hard.

Do not drop the E/O transmitter and O/E receiver, as the impact will possibly damage the inner optical components. 

Do not drop heavy objects like chairs or wheels on the optical fiber as they could possibly damage it.

Please put the optical isolated probe back to its suitcase as delivered. 

Check the optical fiber carefully before use, if there is any damage please stops using immediately.

For better precision, please begin testing after a warm-up of 3~5 minutes.
Power supply connector:5V USB power interface. Please use our company's standard 5V, 2A adapter.

Automatic zero adjustment button: realizes automatic calibration and zero adjustment of output signals.

Automatic calibration indicator: The flashing indicator light indicates that the product is automatically calibrated and zeroed. Two beeps and the light goes out indicate successful calibration; If the indicator light remains on and accompanied by a buzzer for one second after calibration, it indicates calibration failure.

Laser indicator:When the indicator light is on, it indicates that there is laser input, and when the indicator light flashes, it indicates that there is no laser input. If the indicator light is flashing, please check if the front switch is turned on or if the battery is fully charged.
Power indicator: O/E receiver power indicator.

Output connector: standard BNC output connector, connected with standard BNC cable to the oscilloscope.
█ Attenuator
[image: image4.png]


    [image: image5.png]



Attenuator: the attenuator will connect the USB socket of the probe and SSMB RF female socket of the circuit under test. Attenuator have different models with ratio of 20X or 2X, corresponding with range between ±30V / 3V.The standard configuration is 20X.  Users can choose proper attenuator according to the signal under test and set the attenuation parameter on oscilloscope.

SSMB male socket: to gain higher CMRR, users can directly weld the socket to the component or the circuit board of the device under test. The closer SSMB male socket is welded to the connecting line of the component, the higher CMRR will be.

SSMB Male socket to DuPont connector: easy to use in the case of low requirements for CMRR. 

Attenuator (optional)：500X, 2000X attenuators are optional. The corresponding measurement ranges are ± 500V and ± 2500V respectively.
█ Battery
[image: image6.png]Battery position




The product is equipped with two 7.4V/950mAh lithium batteries as standard, as shown in the figure above: battery installation position. In this design, after inserting the attenuator, the attenuator plays a limiting role, and the battery cannot be removed. If you need to remove the battery for charging, refer to the following steps:

First                                         Second
[image: image7.png]Unplug the attenuator plug
—
<

S




[image: image8.png]Push up the cell cover




Third                                         Fourth
[image: image9.png]forwa

Flip the cell cover




      [image: image10.png]Remove the battery as the
inclination angle shown below




 █ Accessories 

[image: image11.jpg]il




     [image: image12.png]-IIHI——-—IIHIIE




Coaxial output cable (CK-310):                   USB cable (CK-315 AM-BM, 1.5m)
1m, connecting oscilloscope and other equipment              O/E receiver power supply cable
[image: image13.png]


                         [image: image14.png]



E/O transmitter support frame (CK-690)
            Battery charger set (CK-691)




(Provide support during test)


   (Including one charger, two batteries and one cable)
             [image: image15.png]


                                [image: image16.png]B




Power supply adaptor (CK-605)*2                     CK-25（BNC/SSMB-JJ）
(5V/2A: for battery charging or receiver power supply)       
[image: image17.png]



CK-322 (SSMB male to RG316 wire approximately 10cm)

Electric Specification
	Model
	OP6031A
	OP6033A
	OP6035A

	Bandwidth(-3dB)
	150MHz
	350MHz
	500MHz

	Rise time
	≤ 2.3ns
	≤ 1.0ns
	≤ 0.7ns

	Typical values of host noise (Vrms)
	2.5mV
	2.5mV
	3mV

	Terminal impedance
	1MΩ

	Output voltage range
	±1.5V

	DC precision
	≤±1%（0~40℃）

	Range Selection (Attenuation Ratio)
	CK-AT2XB
	CK-AT20XB
	CK-AT500XB
	CK-AT2000XB

	Maximum Differential Voltage

under Test  (DC + Peak AC)
	±3V
	±30V
	±500V
	±2500V

	Isolation Voltage  (DC + Peak AC)
	±60kV

	Input Impedance
	2X
	1MΩ//16pF

	
	20X
	1MΩ//6pF

	
	500X
	10MΩ//3pF

	
	2000X
	10MΩ//3pF

	Delay Time
	Host
	About 13ns

	
	BNC(1m)
	5ns

	CMRR Typical Value (Using 20X

plug connector and SSMB socket)
	DC-10MHz
	160dB

	
	10MHz-100MHz
	100dB

	
	100MHz-300MHz
	90dB

	
	300MHz-500MHz
	80dB

	Battery specifications
	Capacity
	7.4V/950mA

	
	Operating Time
	About 8h

	
	Standby time
	About 30 days

	Rear-end Power Supply
	USB 5V/2A

	Auto-zero set function
	YES


Mechanical Specification

	Model
	Parameter

	Probe Size
	Front-end E/O Transmitter
	Around 103*46*34mm

	
	Rear-end O/E Receiver
	Around 125*82*43mm

	Attenuator Length
	Around 200mm

	Cable Length
	Around 1.5m

	BNC Output Cable(CK-310)
	Around 1m

	Probe Weight
	Around 520g


Environmental Specification

	Model
	Parameter

	Operating Temperature
	0℃～50℃

	Storage Temperature
	-30℃～70℃

	Operating Humidity 
	≤85%RH

	Storage Humidity
	≤90%RH

	Operating Altitude
	3000m

	Storage Altitude
	12000m


Operating Process

Attention:

1. Please use our company's standard adapter for power supply (5V, 2A).

User needs to estimate the range of voltage under test and plug-in the proper attenuator before operating. 

Insert the standard 5V 2A adapter of our company into the USB power supply port of the receiving end, and automatically calibrate and zero.(About 20 seconds)
Connect the BNC signal cable between rear-end of the probe and oscilloscope, setup attenuation ratio according to the plug. Adjust the sensitivity of the oscilloscope according to voltage under test.

Connect the SSMB plug of the attenuator with the SSMB socket of the circuit under test.

Turn on the power supply of the circuit under test and begin testing.

Put the front-end box of the probe overhead if possible to decrease the interference from high-voltage pulse circuit.

Because the front-end of the probe is directly connected with the high voltage of the circuit under test, please switch off the power supply of the circuit under test first before touching the probe.

To save the battery power, turn off the switch when you don’t need it.

Please charge in time if the power indicator of the front-end box turns yellow.

The typical connection diagram is shown below:

[image: image18.png]Ve /
T Circuit under test




Important Tips:

 SHAPE  \* MERGEFORMAT 



Do not touch the front-end box of the probe when measuring signals with high floating ground voltage.

Maintenance

Keep the probe clean and dry.
Clean the probe with a soft and dry cloth if needed, do not use chemicals to clean the probe.
Put the probe back to original case when it’s not needed and keep it in a cool, clean and dry place.
When transporting the probe, be sure to put it into the shockproof packaging provided by our company.
Do not pull the input or output cable hard to avoid damage.
Warranty

Please refer to the warranty card instruction.

Packaging list

	Name
	Number
	Name
	Number

	Voltage Probe
	1
	Battery charger kit (CK-691)
	1

	20X Attenuator (CK-AT20XB)
	1
	O/E transmitter support frame  (CK-690)
	1

	SSMB male socket to DuPont 

cable connector (CK-321)
	2
	BNC/SSMB-JJ(CK-25)
	1

	SSMB male socket (CK-23)
	10
	SSMB(M)/RG316(CK-322)
	1

	BNC output cable(CK-310)
	1
	Instruction manual
	1

	USB power supply cable（CK-315 

AM-BM 1.5m）
	1
	Warranty card
	1

	Power supply adaptor 5V/2A(CK-605)
	2
	Calibration  report
	1


	
	CK-AT2XB
	CK-AT500XB
	CK-AT2000XB

	CK-201(2.54_2p)
	-
	5
	-

	CK-202（5.08_2p）
	-
	-
	5


Packing list for optional attenuators

CYBERTEK 
   SHENZHEN ZHIYONG ELECTRONICS CO., LTD.
Addr：Room A1702, Building 4, TianAn Cyber Park, HuangGe Road, LongGang                    District, ShenZhen City, China

Tel:   （86）400 852 0005 / （86 755）86628000

Q Q:  400 852 0005

Email: cybertek@cybertek.cn                          © Zhiyong Electronics, 2025
Url:   http://www.cybertek.cn                      Published in China, Jun. 1, 2025
This symbol means it is harmful to the machine and human body; you must strictly follow the instruction manual to operate. 








In the case of wrong operation, the user risk injury. The content under this mark records the relevant matters needing attention to avoid such dangers.











 





Warning=n





Notice





The user may suffer minor injuries and material damage from the wrong operation. To avoid such situation, the matters under this mark need attention.





Note





This symbolizes important note about how to use the machine.








NOTE





Warning








1
www.cybertek.cn

